## Slide 1: Increase in DLFOV by correction for coma aberrations
Richard Hills
Cavendish Lab
Cambridge, UK

## Slide 2
First make Zemax versions of the Open Dragone and Crossed Dragone cases.
These used conics, i.e. paraboloidal primaries and ellipsoidal secondaries.
Then make matching versions using polynomials in x and y, up to 6th order.
Here are the Open models:
  From Shaul Zemax conics Zemax polynomials

## Slide 3
And these are the Crossed Dragone versions
The one sent by Shaul >>>
The Zemax versions below:
Conics Polynomials

## Slide 4
With 6th order polynomials the performance is essentially identical to that of the conics. The wavefront errors on axis are < 1 micron.
Here are spot diagrams for fields at +/- 7.5 deg for the Open case. (Note that the “exaggeration” is 5, i.e. spread of spot is enlarged x5 w.r.t. position)
Conics Polynomials

## Slide 5
Optimization in Zemax
Although it is possible to derive the changes to the polynomial terms needed to correct for coma analytically, it is easier to brute force it using the optimization feature in Zemax.
It is simplest to allow all the terms to vary, but this means that one has to introduce additional constraints to keep the gross optical properties to be what we want. Here are I forced the angles of incidence of the edge rays (in +y, -y and x directions) to be fixed at the values of the classical design.
I used the center plus a ring of points at 3.5 degree radius for the Open case and 5 deg for the Crossed case and minimized the rms spot size.
Before and after spot diagrams for the Crossed case.
As expected, coma is removed but astigmatism and curvature of the focal plane remain.

## Slide 6
Plots of the rms wavefront error as a function of angle (degrees) in the FOV for the Open case with and without correction. +/-10deg in X and +/-7.5 in Y
Uncorrected (i.e. parab/hyperb) Corrected
The contours are set at 25, 50, 150, 225, 450 and 900 microns. These correspond to 80% Strehl at 900, 450, 150, 100, 50 and 25GHz. The 10dB edge taper is taken into account in these estimates.
As expected the largest improvement is at the highest frequencies – coma is linear in field angle so it dominates close in. Astigmatism and curvature are quadratic, so they dominate at larger angles, and they are not removed.

## Slide 7
Here are the equivalent plots for the Crossed case.
Uncorrected (i.e. parab/hyperb) Corrected
The contours are still at 25, 50, 150, 225, 450 and 900 microns but the outer 2 no longer make it on to the plots, i.e. the whole field has > 80% Strehl at 50GHz. I used the suggested 25dB taper (which does increase the DLFOV).
Again the main improvement is in the central region. The DLFOV is much larger than for the Open case, but it has to be because the detectors cannot be packed so closely (even in angular terms) if you want this strong a taper.

## Slide 8
Finally I checked on how much of the residual loss of gain is due to field curvature. The answer is quite a lot in the Open case.
Corrected flat focal plane Corrected with convex radius 1.8m
The main effect is that the DLFOV field gets quite a lot wider in the x-direction. This can be understood when one take account of Dragone’s comment that, for off-axis systems, one can only remove astigmatism in one direction (e.g. x or y, not both) and the design he gives is for the case of no astigmatism in x.
With astigmatism removed by the geometry, coma corrected by the shape correction and now curvature removed, there are NO Seidel aberrations left!